## Slide 1: SAFT Techniques Review
Bar Weiss

## Slide 2: Agenda
Derivative smoothening
  Results comparison
  Arch summation effect
SNR analysis
  Resolution problem
  Simulation results

## Slide 3: Sinogram Algorithm – Inputs and Outputs
Inputs
  image representing the optical absorption
  Transducer diameter
  Focal length
Output: sinogram

## Slide 4: SIR in the Paper
Analysis done for a case where the transducer is transmitting an impulse
The time derivative of the SIR represents the pressure wave in any time and any location
Marcel Aditi et al, “Transient Fields of Concave Annular Array”, Academic Press 1981

## Slide 5: Interpretation for Receiving Mode
In receiving mode, the SIR represent the receiver sensitivity
is the signal that will be received if the point generates an impulse
transducer
Point sources

## Slide 6: Time Derivative Computation
In order to avoid numerical errors in SAFT summation we smoothen the derivative operation using a gaussian approximation.
Each point is convolved with the derivative of a gaussian instead of a simple numerical derivative.

## Slide 7: Time Derivative Computation
The new derivative can be considered as convolving the old derivative with a gaussian.
The width of the gaussian chosen is .
HPF
LPF
BPF

## Slide 8: Sinogram Results – Simulation Conditions
Parameter | Notation | Unit | Value
Sensor diameter |  | mm | 11
Focal length |  | mm | 12.3
Speed of sound |  | m/s | 1480
Time Oversampling factor | - | - | 5
Image width |  | mm | 4.1
Spatial resolution |  | Pixels | 250
Detector number | - | - | 250
Space between points |  | mm | 0.416
Parameter | Notation | Unit | Value
Sensor diameter |  | mm | 11
Focal length |  | mm | 12.3
Speed of sound |  | m/s | 1480
Time Oversampling factor | - | - | 5
Image width |  | mm | 4.1
Spatial resolution |  | Pixels | 250
Detector number | - | - | 250
Space between points |  | mm | 0.416

## Slide 9: Paper Sinogram

## Slide 10: Sinogram – Simple Derivative
Jake Turner et al “Improved optoacoustic microscopy through three-dimensional spatial impulse response synthetic aperture focusing technique”, Optics Letters, 2014

## Slide 11: SAFT – Simple Derivative

## Slide 12: SAFT Noisy

## Slide 13: Sinogram – Gaussian Derivative

## Slide 14: SAFT – Gaussian Derivative

## Slide 15: Arch summation
Gaussian Derivative
Simple Derivative

## Slide 16: SNR Analysis
We expand the image width to test SAFT’s performance further from the focus.
This raises an issue with the sinogram’s resolution, since when the resolution degrades so does SAFT’s performance.

## Slide 17: Sinogram Results – Simulation Conditions
Parameter | Notation | Unit | Value
Sensor diameter |  | mm | 11
Focal length |  | mm | 12.3
Speed of sound |  | m/s | 1480
Time Oversampling factor | - | - | 5
Image width |  | mm | 12.4
Spatial resolution |  | Pixels | 250/750
Detector number | - | - | 250
Space between points |  | mm | 0.83
Parameter | Notation | Unit | Value
Sensor diameter |  | mm | 11
Focal length |  | mm | 12.3
Speed of sound |  | m/s | 1480
Time Oversampling factor | - | - | 5
Image width |  | mm | 12.4
Spatial resolution |  | Pixels | 250/750
Detector number | - | - | 250
Space between points |  | mm | 0.83

## Slide 18: Noisy Sinogram – % of maxima

## Slide 19: SAFT

## Slide 20: SAFT SIR Noisy

## Slide 21: SNR Comparison

## Slide 22: SNR Paper Results
Jake Turner et al “Improved optoacoustic microscopy through three-dimensional spatial impulse response synthetic aperture focusing technique”, Optics Letters, 2014

## Slide 23: Sinogram Results – Simulation Conditions
Parameter | Notation | Unit | Value
Sensor diameter |  | mm | 11
Focal length |  | mm | 12.3
Speed of sound |  | m/s | 1480
Time Oversampling factor | - | - | 5
Image width |  | mm | 41
Spatial resolution |  | Pixels | 250/750
Detector number | - | - | 250
Space between points |  | mm | 2.8
Parameter | Notation | Unit | Value
Sensor diameter |  | mm | 11
Focal length |  | mm | 12.3
Speed of sound |  | m/s | 1480
Time Oversampling factor | - | - | 5
Image width |  | mm | 41
Spatial resolution |  | Pixels | 250/750
Detector number | - | - | 250
Space between points |  | mm | 2.8

## Slide 24: Noisy Sinogram – % of maxima

## Slide 25: SAFT

## Slide 26: SAFT SIR Noisy

## Slide 27: SNR Comparison

## Slide 28: Backup

## Slide 29: References
Rongkang Gao, Qiang Xue, Yaguang Ren, Hai Zhang, Liang Song, Chengbo Liu, “Achieving depth-independent lateral resolution in AR-PAM using the synthetic-aperture focusing technique”, Photoacoustics, Volume 26, 2022
Jake Turner, Héctor Estrada, Moritz Kneipp and Daniel Razansky “Improved optoacoustic microscopy through three-dimensional spatial impulse response synthetic aperture focusing technique”, Optics Letters, 2014
Joging Park, Seungwan Jeon, Jing Meng, Liang Song, Jin S. Lee and Chulhong Kim, “delay-muliply-and-sum-based synthetic aperture focusing technique in photoacoustic microscopy”, Journal of Biomedical Optics, 2016

## Slide 30: Basic SAFT
The Basic SAFT algorithm treats all the points the same way

## Slide 31: Basic SAFT – Important points
The summation extent is depth independent for the basic SAFT.
Because the extent of summation is large, as we near the focus more invalid signals are included in the summation.

## Slide 32: Coherence Factor
Coherence factor measures in-phase summation for each point in the SAFT image

## Slide 33: SIR SAFT
SIR – Spatial Impulse Response
This technique includes the shape of the SIR in the summation

## Slide 34: SIR SAFT - another approach
Another way to consider the SIR is a binary weighting map for the summation

## Slide 35: SAFT SIR – Hyperbolic Geometry
This is a way to create the binary weighting map, by considering hyperbolic geometry instead of triangular.
The previous method is a simplification of this.

## Slide 36: DMAS SAFT
DMAS – Delay Multiply And Sum
The idea is to extract coherence information based on spatial cross-correlation

## Slide 37: DMAS SAFT – Why Filter?
The scan lines multiplication translates to convolution in the frequency domain, this creates a strong low frequency component that causes artifacts and should be filtered.

## Slide 38: 2D-SAFT

## Slide 39: Adaptive SAFT
The angle is extracted from depth estimation in a MAP (Maximum Amplitude Projection) image.

## Slide 40: 2D Methods Comparison

## Slide 41: Summary

## Slide 42: SAFT Paper Results

## Slide 43: Paper Simulation Conditions
SIR simulated with Field II software package
Noise with of maxima
Transducer with 25MHz center frequency, 11mm outer diameter, focal length of 12.3mm

## Slide 44: Sinogram for 2mm and 10 mm image

## Slide 45: SIR derivative

## Slide 46: Original SIR – Velocity Potential

## Slide 47: Sinogram for

## Slide 48: SAFT -

## Slide 49: SAFT CF -

## Slide 50: SAFT SIR -

## Slide 51: SAFT SIR CF -

## Slide 52: Noisy Sinogram for

## Slide 53: SAFT Noisy -

## Slide 54: SAFT CF Noisy -

## Slide 55: SAFT SIR Noisy -

## Slide 56: SAFT SIR CF Noisy -

## Slide 57: Sinogram 15% Noise

## Slide 58: SAFT SIR CF 10% Noise

## Slide 59: SAFT Just Delay

## Slide 60: Arch summation

## Slide 61: Arch summation